DOCUMENTATION FOR THE PRC PAM SYSTEM

THERE ARE THREE OVERALL COMPONENTS OF THE PRC PAM SYSTEM

I	PAMILA V1.0

-------------------

THIS IS A FOCAL PROGRAM (MODIFIED FROM THE DIGITAL EQUIP. CORP.

SUPPLIED PROGRAM) RUNNING ON THE PDP 8/E WHICH CAN

EITHER OUTPUT CALCULATED DATA TO THE TTY OR SEND IT TO THE

PDP 11/40-RSTS AFTER THE COMPLETION OF EACH RUN.  THE ORDER

OF NUMBERS SENT TO RSTS IS AS FOLLOWS:##

#	   DESCRIPTION

0	   INST #

1	   DAY OF THE YEAR

2	   MIN OF THE DAY

3	   # OF PEAKS

4	   METHOD #

5	   ANAL CODE

6	   1ST PEAK #   OR:  -3 IF NO PEAKS; -4 IF NO AREA

-11 IF NO STANDARD

7	   1ST PEAK AREA OR CONC

8	   1ST PEAK RETENTION TIME

9	   1ST PEAK HEIGHT

10	   1ST PEAK TYPE:  -5 IF BB; -6 IF VB; -7 IF VV; -8 IF BV

.

.

.

5*(N-1)+6  LAST PEAK #

5*(N-1)+7  LAST PEAK CONC

5*(N-1)+8  LAST PEAK RET. T.

5*(N-1)+9  LAST PEAK HEIGHT

5*(N-1)+10 LAST PEAK TYPE

5*(N-1)+11 TOTAL AREA

5*(N-1)+12 -10  : END OF DATA

THERE IS CA. A 3 SEC DELAY AFTER SENDING THE INFO FOR EACH PEAK

AND CA. A 10 SEC DELAY BETWEEN RUNS TO ALLOW RSTS A CHANCE TO KEEP UP

DURING PERIODS OF HEAVY USAGE.

II	PAM12

-----------------

THIS IS A BASIC PROGRAM RUNNING IN DETACHED MODE UNDER ACCOUNT

[1,2] OF RSTS (PDP-11/40).  IT RECEIVES THE NUMBERS SENT FROM

THE PDP-8/E AND WRITES THEM INTO A VIRTUAL CORE FILE NAMED

"PAMNEW".  NORMALLY THE FILES ARE ON DK1:.  HOWEVER, WHEN DK1: IS

UNAVIABLE DK0: IS USED.  THE ACTUAL ACCOUNT USED IS DETERMINED BY THE

INSTRUMENT NUMBER (CHANNEL).  THE STRUCTURE OF EACH FILE IS AS

FOLLOWS:

P1%(0),R1%(200),B(15000) , WHERE

P1%(0) IS THE NEXT LOCATION TO WRITE IN B

R1%(0) IS THE LAST RUN NUMBER WRITTEN

R1%(1-200) ARE THE LOCATIONS IN B FOR THE

START OF THE RESPECTIVE RUNS

III	PAM, PAM1, PAM2

-----------------------------

THESE ARE THREE COMPANION PROGRAMS IN THE RSTS LIBRARY WHICH

ACT ON THE "PAMNEW" FILES AND ASSOCIATED DATA FILES

ACCORDING TO THE OPTIONS LISTED BELOW.  PROGRAMS PAMSTD AND PAMCAL ARE

CHAINED EXTENSIONS OF OPTIONS 11,12, AND 15.

NOTE:  COMMENTS APPEAR FOLLOWING THE "\".

ACTUAL COMPUTER/USER DIALOGUE OUTPUT IS USED FOR CLARITY.

RUN $PAM	\GET PROGRAM FROM RSTS LIBRARY AND RUN

PAM  V04A-2	\IDENTIFING HEADER WITH VERSION NUMBER

OPTION?		\WHEN OPTIONS ARE NOT KNOWN, TYPE <CR> FOR LIST

PAM  V04A-2  OPTIONS ARE:

1	PRINT ALL GC DATA						\PAM

2	PRINT ALL GC DATA & NAME PEAKS USING SPECIFIED WINDOW FILE	\PAM

3	UPDATE WINDOW FILE USING SPECIFIED RUN #			\PAM

4	TRANSFER SPECIFIED RUNS IN ASCII FORMAT (EG TO DEC TAPE)	\PAM

5	OPTION 1 WITH GROUP PEAK					\PAM

6	UNUSED								\PAM

7	UNUSED								\PAM

8	UNUSED								\PAM

9	PRINT ABREVIATED GC DATA FOR SPECIFIED R.T. WINDOW		\PAM1

10	PRINT RUN HEADERS ONLY						\PAM1

11	ASSIGN SAMPLE #'S TO INDIVIDUAL RUN #'S (INITIATE REPORT)	\PAM1

12	MANUAL INPUT OF SAMPLE #'S AND AREAS (INITIATE REPORT)		\PAM1

13	CHANGE RUNS MATRIX DATA						\PAM1

14	OPTION 9 USING ANY ONE COMPOUND FROM WINDOW FILE		\PAM1

15	OPTION 11 WITH SUMMED GROUP PEAK				\PAM1

16	UNUSED								\PAM1

17	CREATE A WINDOW FILE						\PAM2

18	PRINT STATUS OF ALL WINDOW FILES				\PAM2

19	LIST CONTENTS OF SPECIFIC WINDOW FILE				\PAM2

20	TRANSFER ASCII FORMATTED RUNS DATA TO DISK RUNS MATRIX		\PAM2

21	CURRENT RUNS MATRIX FILE STATUS					\PAM2

22	ZERO RUNS MATRIX FILE						\PAM2

23	CREATE ASCII FILE OF ALL WINDOW FILES IN NUMERICAL ORDER (PRIV)	\PAM2

24	REBUILD WINDOW FILES FROM 23'S ASCII FILE (PRIV)		\PAM2

25	UNUSED								\PAM2

\OPTION 1:  PRINT ALL GC DATA

OPTION? 1	\QUESTION REASKED AND USER RESPONDED OPTION 1

\OPTION QUESTION FORMAT IS:

\OPTIONAL DEVICE OR FILE<OPTIONAL DISK:OPTION #.  THUS,

\OPTION? [1,5]DK1:PAMDUP<40><DK0:2, WOULD WRITE THE

\OUTPUT OF OPTION 2, USING DATA FILES STORED ON DISK 0,

\INTO FILE "PAMDUP" WITH PROTECTION CODE 40 ON DISK 1 OF

\ACCOUNT 1,5.  A DEFAULT ANSWER, IE. OPTION? 2, MEANS

\USE KB: (YOUR KEYBOARD) AS THE OUTPUT DEVICE AND WORK WITH

\DATA FILES STORED ON DISK 1 (DK1:)

\ONLY OPTIONS 1,2,9-12,14, AND 15 PRESENTLY CAN USE AN

\OPTIONAL OUTPUT DEVICE OR FILE

\

WHICH CHAN OR ACCT? 10,3	\ONLY ASKED IN  [1,X] ACCOUNTS

\COULD ALSO BE ANSWERED AS "3" FOR CHANNEL 3 OR AS "[10,3]"

\

LAST RUN # WAS 134		\THERE ARE 134 RUNS

\IF LAST RUN IS 200 OR IF STORAGE SPACE HAS BEEN USED UP THE

\MESSAGE "ACCT FULL.  CHECK BCK FILE" IS PRINTED TO REMIND USER THAT

\ANY RUN OVERFLOW IS IN THE '[1,2]DK1:PAMNEW' BACKUP FILE WHICH

\CAN BE READ ONLY FROM [1,X] ACCOUNTS

\

WHICH RUNS DO YOU WANT? 3	\3 = RUN 3 ONLY

\3,6 = RUNS 3 AND 6 ONLY

\5-10,2 = RUNS 5 THRU 10 AND 2

\120- = RUNS 120 THRU 134 (LAST RUN)

\DEFAULT, IE <CR>, SETS UP LAST RUN ONLY

\ANY INCORRECT ANSWER RESULTS IN THE ERROR MESSAGE

\"IMPROPER RUNS SPEC" AND THE PROGRAM MUST BE RERUN

RUN # 3     24-Nov-75  06:45 PM   CHAN 1    METH 1    AC 0   PEAKS 13

PEAK       AREA    RET TIME      HEIGHT    TYPE

1       4,800       1.008       1,252     BB

2     116,333       1.968      14,189     BB

3     136,587       2.595      13,533     BB

4     130,830       3.148      11,662     BB

5     159,156       3.822      12,242     BB

6      99,213       5.082       7,486     BV

7     177,485       5.402      10,336     VB

8     204,611       7.455       9,529     BV

9     168,775       8.088       7,553     VB

10      98,025       9.768       3,924     BV

11     274,131      11.060       7,553     VB

12     148,134      13.170       4,789     BB

13     185,800      23.060       2,773     BB

TOTAL AREA = .190388E 7

READY

\OPTION 2:  PRINT ALL GC DATA & NAME PEAKS USING SPECIFIED WINDOW FILE

RUNNH

PAM  V04A-2

OPTION? 2

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUNS DO YOU WANT? 3

WINDOW FILE #? 21	\WHICH WINDOW FILE OF RETENTION TIMES, NAMES

\IS DATA TO BE COMPARED TO?

\THERE ARE TWO LIBRARIES OF WINDOW FILES

\(DK1: & DK0:).  THE DEFAULT DATA DISK IS

\ASSUMED UNLESS OVERRIDDEN.

\EG. WINDOW FILE? DK0:21

\ONLY OPTIONS 2,3,11, & 12 CAN BE THUSLY OVERRIDDEN

WINDOW ID:  MILL CREEK :CL-HC 9/8/75 STD GC1, LAST UPDATED 10-Dec-75

RUN # 3     24-Nov-75  06:45 PM   CHAN 1    METH 1    AC 0   PEAKS 13

PEAK       AREA    RET TIME      HEIGHT    TYPE NAME

1       4,800       1.008       1,252     BB

2     116,333       1.968      14,189     BB  HEXA-CL-BENZENE

3     136,587       2.595      13,533     BB  LINDANE

4     130,830       3.148      11,662     BB  HEPTACHLOR

5     159,156       3.822      12,242     BB  ALDRIN

6      99,213       5.082       7,486     BV  KELTHANE

7     177,485       5.402      10,336     VB  HEPTACHLOR EPOX

8     204,611       7.455       9,529     BV  DDE-P,P'

9     168,775       8.088       7,553     VB  DIELDRIN

10      98,025       9.768       3,924     BV  ENDRIN

11     274,131      11.060       7,553     VB  DDD-P,P'

12     148,134      13.170       4,789     BB  DDT-P,P'

13     185,800      23.060       2,773     BB  METHOXYCHLOR

TOTAL AREA = .190388E 7

READY

\OPTION 3:  UPDATE WINDOW FILE USING SPECIFIED RUN #

RUNNH

PAM  V04A-2

OPTION? 3

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUN DO YOU WANT? 3	\ONLY SPECIFY ONE RUN!

WINDOW FILE #? 21

WINDOW ID:  MILL CREEK :CL-HC 9/8/75 STD GC1, LAST UPDATED 10-Dec-75

TOLERANCE IS 3 %.  TYPE NEW VALUE TO CHANGE ELSE <CR>?	\NO CHANGE MADE

RUN # 3     24-Nov-75  06:45 PM   CHAN 1    METH 1    AC 0   PEAKS 13

PEAK       AREA    RET TIME      HEIGHT    TYPE NAME        OLD RT  UPDT

1       4,800       1.008       1,252     BB

2     116,333       1.968      14,189     BB  HEXA-CL-BENZ  2.02 ? $   \NO

3     136,587       2.595      13,533     BB  LINDANE       2.63 ? $

4     130,830       3.148      11,662     BB  HEPTACHLOR    3.19 ? Y$  \YES

5     159,156       3.822      12,242     BB  ALDRIN        3.85 ? Y$

6      99,213       5.082       7,486     BV  KELTHANE      5.10 ? Y$

7     177,485       5.402      10,336     VB  HEPTACHLOR E  5.43 ? Y$

8     204,611       7.455       9,529     BV  DDE-P,P'      7.50 ? Y$

9     168,775       8.088       7,553     VB  DIELDRIN      8.17 ? Y$

10      98,025       9.768       3,924     BV  ENDRIN        9.87 ? Y$

11     274,131      11.060       7,553     VB  DDD-P,P'     11.05 ? Y$

12     148,134      13.170       4,789     BB  DDT-P,P'     13.17 ? Y$

13     185,800      23.060       2,773     BB  METHOXYCHLOR 23.25 ? Y$

TOTAL AREA = .190388E 7

\"$" IN UPDATE COLUMN IS <ALT> OR <ESC>

ANY WINDOWS TO REASSIGN? Y	\USED MAINLY WHEN CPD IS

\OUTSIDE OF TOLERANCE AND IS NOT LISTED

TYPE <CR> IF RET TIME IS OK ELSE NEW VALUE

HEXA-CL-BENZENE     2.017	? 1.968	\CHANGE

LINDANE    	    2.630	? 2.595

HEPTACHLOR	    3.148	?	\LEAVE THE SAME

ALDRIN		    3.822	?

KELTHANE	    5.082	?

HEPTACHLOR EPOX	    5.402	?

DDE-P,P'	    7.455	?

DIELDRIN	    8.088	?

ENDRIN		    9.768	?

DDD-P,P'	   11.060	?

DDT-P,P'	   13.170	?

METHOXYCHLOR	   23.060	?

\OPTION 3 CONT

ANY REL CONCS TO CHANGE? Y	\MULTIPLE STDS DON'T HAVE

\TO HAVE THE SAME CONC

TYPE <CR> IF REL CONC IS OK ELSE NEW VALUE

HEXA-CL-BENZENE      1.000	?	\LEAVE THE SAME

LINDANE              1.000	?

HEPTACHLOR	     1.000	?

ALDRIN		     1.000	?

KELTANE		     2.000	?

HEPTACHLOR EPOX	     1.000	?

DDE-P,P'	     1.000	?

DIELDRIN	     1.000	?

ENDRIN	             1.000	?

DDD-P,P'	     1.000	?

DDT-P,P'	     3.000	?

METHOXYCHLOR	     4.000	?

READY

\OPTION 4:  TRANSFER SPECIFIED RUNS IN ASCII FORMAT (EG. TO DEC TAPE)

RUNNH

PAM  V04A-2

OPTION? 4

TAKE FROM WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUNS DO YOU WANT? 1-

OUTPUT FILE NAME? DT1:751216.MEZ	\PUT FILE ON DECTAPE UNIT 1

\UNDER FILE 751216.MEZ

\OUTPUT DEVICE CAN BE ANY

\LEGAL DEVICE. IE. KB:,DK1:, ETC

READY		\TRANSFER COMPLETE

\OPTION 5:  OPTION 1 WITH GROUP PEAK

RUNNH

PAM  V04A-2

OPTION? 5

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCTS

LAST RUN # WAS 23

WHICH RUNS DO YOU WANT? 2

INPUT LOW, HIGH RET TIMES FOR GROUP OF PEAKS? 3,8

RUN # 2     21-Nov-75  05:41 PM   CHAN 1    METH 1    AC 0   PEAKS 13

PEAK       AREA    RET TIME      HEIGHT    TYPE

1     302,417       0.320      22,820     BB

2      66,698       1.958       9,275     BB

3      86,674       2.578       9,016     BB

4*     89,626       3.142       7,934     BB

5*    108,350       3.775       8,076     BB

6*     62,421       5.035       4,498     BV

7*    121,402       5.362       6,688     VB

8*    124,671       7.428       5,811     BV

*TOT    506,470                  33,006    GROUP

9     105,994       8.055       4,744     VB

10      77,255       9.655       2,227     BV

11     163,629      11.070       4,551     VB

12      89,388      13.040       2,322     BB

13      88,788      23.070       1,453     BB

TOTAL AREA = .148731E 7

\OPTIONS 6-8 ARE UNUSED AT PRESENT

\OPTION 9:  PRINT ABREVIATED GC DATA FOR SPECIFIED R.T. WINDOW

RUNNH

PAM  V04A-2

OPTION? 9

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUNS DO YOU WANT? 1-10

INPUT LOW,HIGH RET TIMES TO DEFINE WINDOW? 3,3.2	\LOW FIRST!!!

COMPOUND -- HEPTACHLOR	\USER TYPES IDENTIFER

RUN CHAN   DATE      TIME   RET TIME        AREA       HEIGHT

1   1  24-Nov-75 05:32 PM    3.118      22,680        1,838

2   1  24-Nov-75 06:08 PM    3.152      76,769        7,379

3   1  24-Nov-75 06:45 PM    3.148     130,830       11,662

4   1  24-Nov-75 07:21 PM    3.160     316,110       20,090

5   1  24-Nov-75 07:58 PM     NONE

6   1  24-Nov-75 08:35 PM    3.160      17,629        1,123

7   1  24-Nov-75 09:11 PM     NONE

8   1  24-Nov-75 09:47 PM     NONE

9   1  24-Nov-75 10:24 PM     NONE

10   1  24-Nov-75 11:00 PM    3.180      10,289          750

\OPTION 10:  PRINT RUN HEADERS ONLY

RUNNH

PAM  V04A-2

OPTION? 10

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUNS DO YOU WANT? 1-10

RUN CHAN   DATE      TIME   PEAKS

1   1  24-Nov-75 05:32 PM  2

2   1  24-Nov-75 06:08 PM  8

3   1  24-Nov-75 06:45 PM  12

4   1  24-Nov-75 07:21 PM  11

5   1  24-Nov-75 07:58 PM  13

6   1  24-Nov-75 08:35 PM  0

7   1  24-Nov-75 09:11 PM  4

8   1  24-Nov-75 09:47 PM  5

9   1  24-Nov-75 10:24 PM  12

10   1  24-Nov-75 11:00 PM  8

READY

\OPTION 11:  ASSIGN SAMPLE #'S TO INDIVIDUAL RUN #'S (INITIATE REPORT)

RUNNH

PAM  V04A-2

OPTION? 11

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 134

WHICH RUNS DO YOU WANT? 1-9

WINDOW FILE #? 10

WINDOW ID:  MILL CREEK :CL-HC DEMO, LAST UPDATED 10-Dec-75

USE PEAK AREAS(0) OR HGTS(1)? 0	\CALCULATION CAN BE BASED

\ON EITHER AREA OR HEIGHT

EXT STD WILL ASSUME SAMPLE #'S<101 TO BE IN NG INJ

RUN CHAN   DATE      TIME    SAMPLE #

1   1  24-Nov-75 05:32 PM? .1

2   1  24-Nov-75 06:08 PM? .5

3   1  24-Nov-75 06:45 PM? 1

4   1  24-Nov-75 07:21 PM? 2

5   1  24-Nov-75 07:58 PM? 	\SAMPLES MAY BE SKIPPED BY TYPING ONLY <CR>

6   1  24-Nov-75 08:35 PM? 5023

7   1  24-Nov-75 09:11 PM? 5024

8   1  24-Nov-75 09:47 PM? 5025

9   1  24-Nov-75 10:24 PM? 5026

WANT EXT STD CURVE FIT? Y	\OPTION TO HAVE STANDARD CURVES

\CALCULATED OF THE FORM A + BX + CX^2

\FOR EACH STANDARD IN WINDOW FILE.

\PROGRAM CHAINS TO "$PAMSTD"

STD CURVE(S) COMPLETE

WANT TO GRAPH ANY OF THE STD CURVES? Y	\CHAINS TO "$FGRAPH" 20000

WHICH STD # DO YOU WANT?	\DEFAULT ANSWER OR A NUMBER OUTSIDE THE RANGE

\GIVES LIST OF AVAILABLE STDS AND THEIR #'S

1            HEXA-CL-BENZENE

2            SIMAZINE

3            LINDANE

4            HEPTACHLOR

5            ALDRIN

6            KELTHANE

7            HEPTACHLOR EPOX

8            DDE-P,P'

9            DIELDRIN

10           ENDRIN

11           DDD-P,P'

12           DDT-P,P'

13           METHOXYCHLOR

WHICH STD # DO YOU WANT? 2

INSUFFICIENT POINTS FOR STD CURVE OF SIMAZINE	\"$PAMSTD" DIDN'T CALC CURVE

WHICH STD # DO YOU WANT? 11

BEGINNING AND END POINTS FOR THE GRAPH.  DEFAULT VALUES ARE 0 AND 82.0736 ?

\THE DEFAULT VALUES ARE THE MIN AND MAX FOR THE STD CHOSEN.

\OTHER END POINTS MAY BE SPECIFIED.  E.G.  ? 0,30

\

OUTPUT FILE NAME (DEFAULT IS YOUR KEYBOARD)?	\ONE COULD OUTPUT THE GRAPH

\TO THE LINE PRINTER (WHEN

\AVAILABLE) OR WRITE THE GRAPH

\INTO A DISK FILE FOR LATER

\DISPLAY

LLOP? Y	\GET TO LAST LINE OF PAGE FOR DECWRITER

STD CURVE FOR DDD-P,P'; X-AXIS/10^ 4 .  Y-AXIS IN NG INJ

-.5         0        .5         1       1.5         2       2.5         3

0       +YYYYYYYYYOYYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+

X         +*        +         +         +         +         +         X

X         +  O      +         +         +         +         +         X

X         +    *    +         +         +         +         +         X

X         +     *   +         +         +         +         +         X

8.20735 +         +       * +         +         +         +         +         +

X         +        *+         +         +         +         +         X

X         +         +*        +         +         +         +         X

X         +         + *       +         +         +         +         X

X         +         +   *     +         +         +         +         X

16.4147 +         +         +  O *    +         +         +         +         +

X         +         +      *  +         +         +         +         X

X         +         +       * +         +         +         +         X

X         +         +         *         +         +         +         X

X         +         +         +*        +         +         +         X

24.6221 +         +         +         + *       +         +         +         +

X         +         +         +   *     +         +         +         X

X         +         +         +    *O   +         +         +         X

X         +         +         +     *   +         +         +         X

X         +         +         +      *  +         +         +         X

32.8294 +         +         +         +        *+         +         +         +

X         +         +         +         *         +         +         X

X         +         +         +         +*        +         +         X

X         +         +         +         + *       +         +         X

X         +         +         +         +  *      +         +         X

41.0368 +         +         +         +         +   *     +         +         +

X         +         +         +         +    *    +         +         X

X         +         +         +         +     *   +         +         X

X         +         +         +         +      *  +         +         X

X         +         +         +         +       * +         +         X

49.2441 +         +         +         +         +        *+         +         +

X         +         +         +         +         *         +         X

X         +         +         +         +         +*        +         X

X         +         +         +         +         + *       +         X

X         +         +         +         +         +  *      +         X

57.4515 +         +         +         +         +         +  *      +         +

X         +         +         +         +         +   *     +         X

X         +         +         +         +         +    *    +         X

X         +         +         +         +         +     *   +         X

X         +         +         +         +         +     *   +         X

65.6588 +         +         +         +         +         +      *  +         +

X         +         +         +         +         +       * +         X

X         +         +         +         +         +       * +         X

X         +         +         +         +         +        *+         X

X         +         +         +         +         +        *+         X

73.8662 +         +         +         +         +         +         *         +

X         +         +         +         +         +         *         X

X         +         +         +         +         +         +*        X

X         +         +         +         +         +         +*        X

X         +         +         +         +         +         + *       X

82.0736 +YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YOYYYYYYY+

EQUATION COEFS:

X^0  -.247225E-1

X^1   .526024E-1

X^2  -.250611E-3

PLOT DIFFERENT STD? Y		\ALSO ANSWER "Y" IF DIFFERENT X-AXIS

\END POINTS ARE WANTED

WHICH STD # DO YOU WANT? 12

BEGINNING AND END POINTS FOR THE GRAPH.  DEFAULT VALUES ARE 0 AND 65.0494?

OUTPUT FILE NAME (DEFAULT IS YOUR KEYBOARD)?

LLOP? Y

STD CURVE FOR DDT-P,P'; X-AXIS/10^ 4 .  Y-AXIS IN NG INJ

-.5         0        .5         1       1.5         2       2.5         3

0       +YYYYYYYY*OYYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+YYYYYYYYY+

X         *         +         +         +         +         +         X

X         *         +         +         +         +         +         X

X         +*        +         +         +         +         +         X

X         + *       +         +         +         +         +         X

1       +         +  *      +         +         +         +         +         +

X         +   *     +         +         +         +         +         X

X         +    *    +         +         +         +         +         X

X         +    *    +         +         +         +         +         X

X         +     *   +         +         +         +         +         X

2       +         +      *  +         +         +         +         +         +

X         +       * ^C		\CTRL C ABORTS THE GRAPH

PLOT DIFFERENT STD? N

WANT TO SEE CALCULATED DATA? Y	\OPTION FOR DATA REPORT

\PROGRAM CHAINS TO "$PAMCAL" 31000

REPORT ID? DEMO FOR MILL CREEK		\ANY ONE LINE

WANT ANSWERS IN NG, UG/CM^2, PPM, PPB, OR PPT? PPM

HOW MANY DECIMAL PLACES? 3

DETECTION LIMIT? .006		\OPTIONAL, BUT IF SPECIFIED ANY VALUE LESS

\THAN THE LIMIT WILL APPEAR AS "N.D."

IF ALL SAMPLE INJECTION VOLUMNS WERE EQUAL THEN GIVE IN UL ELSE <CR>? 5

IF ALL SAMPLE FINAL VOLUMNS WERE EQUAL THEN GIVE IN ML ELSE <CR>?

IF ALL SAMPLE ORIGINAL WEIGHTS WERE EQUAL THEN GIVE IN GM ELSE <CR>?

\FOR UG/CM^2    AREAS                          CM^2

\ABOVE 3 QUESTIONS NOT ASKED IF UNITS ARE NG

USE <ALT> OR <ESC> TO INPUT ANSWERS	\ONLY ASKS FOR DATA NOT GIVEN ABOVE

SAMPLE #            FINAL VOLUMN, ML     ORIGINAL WEIGHT,GM

5023     	    ? 1$       	         ? 1.23$

5024                ? 2$                 ? 1.17$

5025                ? 1$                 ? 1.09$

5026                ? 3$                 ? 1.01$

LLOP?		\LAST LINE OF PAGE.  REASKED UNTIL ANSWERED "Y"

LLOP? Y		\(USEFUL WITH FANFOLD PAPER)

\OPTION 11 CONT

25-Jan-76

DEMO FOR MILL CREEK

***PPM***

SAMPLE # :             5023          5024          5025          5026

HEXA-CL-BENZENE       0.007         0.062         0.078         1.256

SIMAZINE             TFSTDS        TFSTDS        TFSTDS        TFSTDS

LINDANE               0.006         0.075         0.074         1.258

HEPTACHLOR            0.006         0.082         0.072         1.259

ALDRIN                0.007         0.081         0.072         1.259

KELTHANE              0.009         0.081         0.071         1.260

HEPTACHLOR EPOX       0.010         0.076         0.073         1.259

DDE-P,P'              0.008         0.064         0.076         1.257

DIELDRIN              0.008         0.085         0.070         1.260

ENDRIN                 N.D.         0.088         0.069         1.260

DDD-P,P'              0.008         0.093         0.068         1.261

DDT-P,P'               N.D.         0.090         0.069         1.261

METHOXYCHLOR           N.D.         0.091         0.069         1.261

N.D. = NOT DETECTED; DETECTION LIMIT = .006 PPM

TFSTDS = TOO FEW STANDARDS FOR QUANTITATION

\OTHER POSSIBLE FOOTNOTE:  "*

VALUE OUTSIDE STANDARDIZATION LIMITS"

\"*" AT END OF TABLE VALUE INDICATES THAT PEAK AREA OR HEIGHT

\(AS CHOSEN) WAS GREATER THAN LARGEST STANDARD.  VALUE REPORTED

\USES STANDARD EQUATION BUT MAY NOT BE A FAIR

#\REPRESENTATION OF THE UNKNOWN CONCENTRATION

READY

\TYPING "RUNNH" AT THIS POINT GENERATES THE FOLLOWING REPORT OPTIONS:

RUNNH

OPTION?		\NO OPTION SPECIFIED GIVES LIST

OPTIONS ARE:

1	RELIST OUTPUT	\DECIMAL PLACES AND DETECTION LIMIT MAY BE CHANGED

2	CHANGE CONC UNITS	\IE, FROM PPM TO PPT, ETC

3	CORRECT SAMPLE DATA	\SAMPLE INJ VOL, FINAL VOL, OR WEIGHT

OPTION?

\TO CHANGE ALL VARIABLES TYPE 'CHAIN "$PAMCAL" 31000'

\OPTION 12:  MANUAL INPUT OF SAMPLE #'S AND AREAS (INITIATE REPORT)

RUNNH

PAM  V04A-2

OPTION? 12

WINDOW FILE #? 90	\WILL ONLY WORK WITH FIRST WINDOW IN FILE

WINDOW ID:  BOB'S GUTHION, LAST UPDATED 20-AUG-76

ENTER VALUES WITH <ALT> OR <ESC>	\IE, DON'T USE <CR> AS TERMINATOR

TERMINATE INPUT WITH SAMPLE #=<ALT>

SAMPLE #      AREA

? 1$	      200$

? 2.4$        307$

? 5.6$        488$

? ABC$        345$

? $				\INPUT TERMINATED.

WANT EXT STD CURVE FIT?		\AT THIS POINT PROCEED LIKE OPTION 11

\OPTION 13:  CHANGE RUNS MATRIX DATA

RUNNH

PAM  V04A-2

OPTION? 13

WHICH CHAN OR ACCT? 1		\ONLY ASKED IN [1,X] ACCOUNTS

LAST RUN # WAS 106

WHICH RUN DO YOU WANT? 3	\CAN ONLY ASK FOR ONE RUN AT A TIME

#	ACTION

1	CHANGE AREA

2	CHANGE RET TIME

3	CHANGE HT

4	DELETE PEAK

'PEAK #' AND 'WHICH RUN...' QUESTIONS WILL BE REASKED UNTIL ANSWERED 0

PEAK #? 12

NO SUCH PEAK #		\MUST USE AN ACTUAL PEAK #

PEAK #? 2		\QUESTION REASKED

ACTION #? 1	\PEAK 2'S AREA WILL BE CHANGED

NEW VALUE? 2.078

PEAK #? 5

ACTION #? 4

PEAK #? 0	\DO A DIFFERENT RUN #

WHICH RUN DO YOU WANT? 4

PEAK #? 1

ACTION #? 4

PEAK #? 0

WHICH RUN DO YOU WANT? 0	\ONLY EXIT FROM OPTION 13 THAT

\WILL BE SURE TO CLOSE FILES

READY

\OPTION 14:  OPTION 9 USING ANY ONE COMPOUND FROM WINDOW FILE

RUNNH

PAM  V04A-2

OPTION? 14

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCTS

LAST RUN # WAS 23

WHICH RUNS DO YOU WANT? 5-8

WINDOW FILE #? 10

WINDOW ID:  MILL CREEK:  CL-HC DEMO, LAST UPDATED 09-OCT-76

WHICH COMPOUND # DO YOU WANT FROM WINDOW FILE ?	\DEFAULT OR # OUTSIDE

\RANGE PRODUCES LIST

1		HEXA-CL-BENZENE

2		SIMAZINE

3		LINDANE

4		HEPTACHLOR

5		ALDRIN

6		KELTHANE

7		HEPTACHLOR EPOX

8		DDE-P,P'

9		DIELDRIN

10		ENDRIN

11		DDD-P,P'

12		DDT-P,P'

13		METHOXYCHLOR

WHICH COMPOUND # DO YOU WANT FROM WINDOW FILE? 9

09-Jun-77   SELECTIVE LISTING FOR DIELDRIN

NOMINAL RET. TIME = 8.027 +- 0.241 MIN	\CALC FROM TOLERANCE IN WINDOW FILE

RUN CHAN   DATE      TIME   RET TIME        AREA       HEIGHT

5   1  21-Nov-75 05:05 PM    8.018      20,631          912

6   1  21-Nov-75 05:41 PM    8.055     105,994        4,744

7   1  21-Nov-75 06:17 PM    8.065     185,572        7,797

8   1  21-Nov-75 06:54 PM    8.027     503,559       17,340

READY

\OPTION 15:  OPTION 11 WITH A SUMMED GROUP PEAK

RUNNH

PAM  V04A-2

OPTION? 15

WHICH CHAN OR ACCT? 10,1	\ONLY ASKED IN [1,X] ACCTS

LAST RUN # WAS 23

WHICH RUNS DO YOU WANT? 1-8

LOW,HIGH RET TIMES FOR GROUP PEAK? 3,8

WINDOW FILE #/ 10

WINDOW ID:  MILL CREEK:  CL-HC DEMO, LAST UPDATED 09-OCT-76

USE PEAK AREAS(0) OR HGTS(1)? 0

EXT STD WILL ASSUME SAMPLE #'S<101 TO BE IN NG INJ

RUN CHAN   DATE      TIME   SAMPLE #

1   1  21-NOV-75 05:05 PM?		\PROCEEDS JUST LIKE OPTION 11

\OPTION 16 IS UNUSED AT PRESENT

\OPTION 17:  CREATE A WINDOW FILE

RUNNH

PAM  V04A-2

OPTION? 17

WINDOW FILE #? 5

\THE MESSAGE "WINDOW REQUESTED CAN'T BE CREATED FROM A

\NON-PRIVILEGED ACCOUNT" WILL BE PRINTED IF WINDOW FILE DOESN'T

\ALREADY EXIST AND USER IS LOGGED INTO A NON-PRIVILEGED ACCOUNT

FILE EXISTS!!!

WINDOW ID:  TEST 2, LAST UPDATED 10-AUG-75

REPLACE? Y	\DESTROYS OLD WINDOW FILE 5

\NEW WINDOW FILES CAN ONLY BE CREATED FROM [1,X] ACCOUNTS

NUMBER OF RETENTION TIME WINDOWS UP TO 20? 3

WINDOW TOLERANCE IN %? 3

WINDOW ID:  (UP TO 32 CHARACTERS)  PAM.DOC TEST

ARE ALL COMPONENTS OF STD THE SAME REL CONC? Y	\IF NOT WILL ASK FOR THEM

\AS WINDOWS ARE ASSIGNED

WINDOW # 1 :

RETENTION TIME? 1.01

NAME? COMPOUND 1

WINDOW # 2 :

RETENTION TIME? 2.07

NAME? COMPOUND 2

WINDOW # 3 :

RETENTION TIME? 3.46

NAME? COMPOUND 3

READY

\OPTION 18:  PRINT STATUS OF ALL WINDOW FILES

RUNNH

PAM  V04A-2

OPTION? 18

WINDOW FILE STATUS [1,2]DK1:  10:37 AM  16-Dec-75

FILE #  UPDATED    # WIND  % TOL  WINDOW ID

1       22-Oct-75  12      3      MILL CREEK:  CL-HYDROCARBONS; GC

11      21-Aug-75  2       3      TEST

10      10-Dec-75  19      3      MILL CREEK:  CL-HC DEMO

21      16-Dec-75  12      3      MILL CREEK :CL-HC 9/8/75 STD GC1

31      31-Oct-75  10      3      ATLANTIC SALMON PCBS; GC1

41      18-Nov-75  19      3      PCB FOR BREAD; HIGHTEST 13 PEAKS

51      25-Sep-75  1       3      PCB 6

61      25-Sep-75  1       3      PCB 3

3       09-Dec-75  1       3      TESAR

4       03-Dec-75  1       10     TEST 2

5       16-Dec-75  3       3      PAM.DOC TEST

READY

\OPTION 19:  LIST CONTENTS OF SPECIFIC WINDOW FILE

RUNNH

PAM  V04A-2

OPTION? 19

TYPE WINDOW FILE # FOR LISTING OF FILE ELSE <CR>? 5

I    RET. T.         REL. C.          NAME

1     1.01             1              COMPOUND 1

2     2.07             1              COMPOUND 2

3     3.46             1              COMPOUND 3

\OPTION 20:  TRANSFER ASCII FORMATTED RUNS DATA TO DISK RUNS MATRIX

RUNNH

PAM  V04A-2

OPTION? 20

PUT IN WHICH CHAN OR ACCT? 10,1		\ONLY ASKED IN [1,X] ACCOUNTS

CONFIRM "DK1:[10,1]PAMNEW<...." OPERATION? Y	\YOU CAN STILL CHANGE YOUR

\MIND.  A "Y" WILL DESTROY

\PRESENT CONTENTS OF PAMNEW

INPUT FILE NAME(S)? DT1:751216.MEZ	\GET DATA FROM DEC TAPE 1

\CAN SPECIFY MORE THAN

\ONE INPUT FILE NAME

\(SEPARATED BY A ","),

\WHICH IS USEFUL WHEN

\COMBINING GOOD RUNS INTO

\ONE FILE OR FOR REARRANGING

\RUNS WITHIN THE MATRIX

\POSSIBLE MESSAGES:

\"OLD SYSTEM DATA.  GIVE 4 DIGIT YEAR"

\"CAN ONLY ACCEPT THE FIRST 200 RUNS OR 15000 TOTAL NUMBERS"

READY

\OPTION 21:  CURRENT RUNS MATRIX FILE STATUS

RUNNH

PAM  V04A-2

OPTION? 21

PAMNEW FILE STATUS ON 16-Dec-75 AT 10:34 AM

ACCOUNT  CHAN  LAST RUN #   MATRIX     DATE

10,0       0      22           326   16-Dec-75

10,1       1     134         4,291   15-Dec-75

10,2       2     175        11,026   12-Dec-75

10,3       3       0             0

10,4       4       0             0

10,5       5       0             0

10,6       6       0             0

10,7       7       0             0

10,8       8       0             0

10,9       9     112         2,888   08-Dec-75

10,10     10       0             0

10,11     11       0             0

1,2     BCK       4           227   12-Dec-75

READY

\ABOVE LIST AVAILABLE ONLY FROM [1,X] ACCOUNTS

\OR BY TYPING "PAM/X" ON NON-LOGGED IN TERMINALS,.

\WHERE "/X" IS OPTIONAL AND IF USED WILL

\LOOK UP ONLY CHANNEL X

\OTHERS RECEIVE ONLY THEIR OWN FILE STATUS.

\LIMITS ARE 200 RUNS OR 15,000 NUMBERS IN MATRIX

\OPTION 22:  ZERO RUNS MATRIX FILE

RUNNH

PAM  V04A-2

OPTION? 22

WHICH CHAN OR ACCT? 10,3	\ONLY ASKED IN [1,X] ACCOUNTS

CONFIRM "DK1:[10,3]PAMNEW" ZERO? Y	\CAN STILL CHANGE YOUR MIND

READY				\OPERATION COMPLETE

\OPTION 23:  CREATE ASCII FILE OF ALL WINDOW FILES IN NUMERICAL

ORDER (PRIV)  [ONLY TO BE USED BY SYSTEM MANAGER]

RUNNH

PAM  V04A-2

OPTION? 23

WINDOW STORAGE FILE? WIND1.STO	\MAY TAKE SEVERAL SEC BEFORE

\THIS QUESTION IS ASKED

READY

\OPTION 24:  REBUILD LIBRARY OF WINDOW FILES FROM 23'S ASCII

FILE (PRIV)  [ONLY TO BE USED BY SYSTEM MANAGER]

RUNNH

PAM  V04A-2

OPTION? 24

WINDOW STORAGE FILE? WIND1.STO

READY		\MAY TAKE 20-30 SEC

\OPTION 25 IS UNUSED AT PRESENT

END OF PAMILA.DOC

14-SEP-76

R. LEAVITT